| CHEMTECH PROJECT NUMBER: Q1447 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF022725.M) for 2-Nitrophenol,2-Nitroaniline,2,6-Dinitrotoluene, 3-Nitroaniline, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenol these compounds are passing on Linear Regression and 2,4-Dinitrophenol is passing on Quadratic regression. The Continuous Calibration File ID BF141824.D met the requirements except for 4-Nitroaniline and 4-Nitrophenol. The associate samples have no positive hit for these compounds therefore no corrective action was required. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1475-01MS} with File ID: BF141819.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[52%], 4-Chloroaniline[33%], Fluoranthene[64%] and Pyrene[64%].these compounds did not meet the NJDKQP criteria but met the in-house criteria. The MSD {Q1475-01MSD} with File ID: BF141820.D recoveries met the acceptable requirements except for 3,3-Dichlorobenzidine[57%] and 4-Chloroaniline[35%].these compounds did not meet the NJDKQP criteria but met the in-house criteria. The Blank Spike for {PB166957BS} with File ID: BF141826.D met requirements for all samples except for 3,3-Dichlorobenzidine[53%], 3-Nitroaniline[56%] and 4-Chloroaniline[29%].these compounds did not meet the NJDKQP criteria but met the in-house criteria. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for WC1 which is fail due to the presence of nontargeted hydrocarbons which can be observed by the abnormal chromatogram. Hence no corrective action is required also TR-04-02282025MS and TR-04-02282025MSD has Internal Standards fail due to matrix interference. Hence, no corrective action is required. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date