| CHEMTECH PROJECT NUMBER: Q1456 | MATRIX: Water |
| --- | --- |
| METHOD: 625.1/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BM020525.M) for Pentachlorophenol which are passing on Linear Regression and 2-Nitrophenol, 2,4-Dinitrophenol, 2,4-Dinitrotoluene,4,6-Dinitro-2-methylphenol, and these are passing on Quadratic regression. The Continuous Calibration File ID BM049675.D met the requirements except for 2,4,6-Trichlorophenol,2,4-Dinitrophenol,2,6-Dinitrotoluene,2-Nitrophenol,3,3-Dichlorobenzidine,4,6-Dinitro-2-methylphenol,Benzidine,Benzo(g,h,i)perylene,Dibenzo(a,h)anthracene,Hexachlorocyclopentadiene and Indeno(1,2,3-cd)pyrene .The associate samples have no positive hit for these compounds therefore no corrective action was required. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for EFF-WASTE WATER [2-Fluorophenol - 47%, Phenol-d6 - 29%]. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB166933BS} with File ID: BM049678.D met requirements for all samples except for 4,6-Dinitro-2-methylphenol[131%], Hexachlorocyclopentadiene[290%] . The Blank Spike Duplicate for {PB166933BSD} with File ID: BM049679.D met requirements for all samples except for Hexachlorocyclopentadiene[280%] . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date