284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922

CASE NARRATIVE

### Alliance Technical Group, LLC - Newark

### Project Name: NJ Waste Water PT

### Project # N/A

### Chemtech Project # Q1503

Test Name: Metals Group3,Metals Group4,Metals Group5,Mercury,Silica

A. Number of Samples and Date of Receipt:

8 Water samples were received on 03/05/2025.

B. Parameters:

According to the Chain of Custody document, the following analyses were requested:

Hardness, Calcium, Hardness, Total, Mercury, Metals Group3, Metals Group4, Metals Group5 and Silica. This data package contains results for Metals Group3,Metals Group4,Metals Group5,Mercury,Silica.

C. Analytical Techniques:

The analysis and digestion of Metals Group3,Metals Group4,Metals Group5,Silica was based on 200.7, The analysis and digestion of Metals Group3,Metals Group4,Metals Group5 was based on 200.8 and The analysis and digestion of Mercury was based on 245.1.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

Sample PT-SIO2-WP was diluted due to high concentrations for Dissolved Silica & Sample PT-HG-WP was diluted due to high concentrations for Mercury.

The Blank Spike met requirements for all samples.

The Blank analysis did not indicate the presence of lab contamination.

The Calibration met the requirements.

E. Additional Comments:

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________