| CHEMTECH PROJECT NUMBER: Q1514 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF031025.M) for 2,4-Dinitrophenol and this compound is passing on Linear Regression. The % RSD is greater than 20% in the Initial Calibration (8270-BG030525.M) for 2,6-Dinitrotoluene, 2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenol, Butylbenzlphthalate, these compounds are passing on Linear Regression and 2-Nitrophenol, 2-Nitroaniline, are passing on Quadratic regression. The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1514-03MS} with File ID: BG064066.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[39%], 3-Nitroaniline[37%], 4-Chloroaniline[20%] and Hexachlorocyclopentadiene[168%] , these compound did not meet the NJDKQP criteria but met the in-house criteria therefor no further no corrective action taken. The MSD {Q1514-03MSD} with File ID: BG064067.D recoveries met the acceptable requirements except for 3,3-Dichlorobenzidine[40%], 3-Nitroaniline[40%], 4-Chloroaniline[23%] and Hexachlorocyclopentadiene[161%] , these compound did not meet the NJDKQP criteria but met the in-house criteria therefor no further no corrective action taken. The Blank Spike for {PB167031BS} with File ID: BG064058.D met requirements for all samples except for 3,3-Dichlorobenzidine[53%], 3-Nitroaniline[51%] and 4-Chloroaniline[39%] , these compounds did not meet the NJDKQP criteria but met the in-house criteria. No further corrective action required. The Blank Spike Duplicate met requirements for all samples .The Blank Spike for {PB167027BS} with File ID: BG064060.D met requirements for all samples except for 3-Nitroaniline[51%], 3,3-Dichlorobenzidine[58%], 4-Chloroaniline[36%] , these compounds did not meet the NJDKQP criteria but met the in-house criteria. No further corrective action required and Atrazine[131%], this compound did not meet the NJDKQP criteria and in-house criteria. No further corrective action required. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date