| CHEMTECH PROJECT NUMBER: Q1539 | MATRIX: Water |
| --- | --- |
| METHOD: 6020B,7470A |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (TAPFTA-MW01I-031025-00-T2MS) analysis met criteria for all samples except for Arsenic, Iron, Manganese, Potassium, Silver due to matrix interference. The Matrix Spike Duplicate (TAPFTA-MW01I-031025-00-T2MSD) analysis met criteria for all samples except for Arsenic, Barium, Iron, Manganese, Potassium, Silver due to matrix interference. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria The Duplicate (TAPFTA-MW01I-031025-00-T2DUP) analysis met criteria for all samples except for Aluminum due to matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Tl, Pb, U are being analyzed with Non-Collision Cell. Helium gas is used for the Collision Cell analysis.

_______________________________________				__________________________

QA REVIEW								Date