| CHEMTECH PROJECT NUMBER: Q1572 | MATRIX: Solid |
| --- | --- |
| METHOD: 8082A/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. The Initial Calibration met the requirements . The Continuous Calibration File ID PO109890.D met the requirements except for Decachlorobiphenyl is failing in 2nd column, The Continuous Calibration File ID PO109904.D met the requirements except for Decachlorobiphenyl is failing in 1st column, The Continuous Calibration File ID PO109923.D met the requirements except for Decachlorobiphenyl is failing in 2nd column, The Continuous Calibration File ID PO109964.D met the requirements except for Decachlorobiphenyl is failing in 2nd column, above mentioned CCAL failing in one column but another column passing for that therefore no corrective action taken. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for C0AC9 [Decachlorobiphenyl(2) - 250%], C0AC9DL [Decachlorobiphenyl(2) - 347%], C0O16 [Decachlorobiphenyl(2) - 229%], C0O16DL [Decachlorobiphenyl(2) - 270%], C0O17 [Decachlorobiphenyl(2) - 295%], C0O17DL [Decachlorobiphenyl(2) - 483%], C0O18DL [Decachlorobiphenyl(2) - 178%],C0O20 [Decachlorobiphenyl(2) - 232%], C0O20DL [Decachlorobiphenyl(2) - 311%],AS per method one surrogate allowed to fail to meet the criteria per column. No further corrective action was taken while for, Samples C0O14DL [Decachlorobiphenyl(2) - 232%, Tetrachloro-m-xylene(1) - 150%, Tetrachloro-m-xylene(2) - 167%], C0O15DL2 [Decachlorobiphenyl(2) - 190%, Tetrachloro-m-xylene(2) - 159%], C0O22 [Decachlorobiphenyl(1) - 1865%, Decachlorobiphenyl(2) - 1454%], C0O23 [Decachlorobiphenyl(1) - 222%, Decachlorobiphenyl(2) - 211%], C0O23DL [Decachlorobiphenyl(1) - 262%, Decachlorobiphenyl(2) - 272%], C0O24 [Decachlorobiphenyl(1) - 249%, Decachlorobiphenyl(2) - 349%], these samples required dilution as well due to high concentration therefore samples were reanalyzed with dilution and reported while for, Samples C0AC9DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O14DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O16DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O17DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O18DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O21DL [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O22DL [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O24DL [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], C0O24DL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0% and Tetrachloro-m-xylene(2) - 0%] surrogates were diluted out due to high dilution, no further corrective action was taken. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1572-06MS} with File ID: PP070564.D recoveries met the requirements for all compounds except for AR1016[7313%], AR1260[169%] due to sample matrix interference. The MSD {Q1572-06MSD} with File ID: PP070565.D recoveries met the acceptable requirements except for AR1016[7777%], AR1260[215%] due to sample matrix interference. The Blank Spike met requirements for all samples . The RPD for {Q1572-06MSD} with File ID: PP070565.D met criteria except for AR1260[24%] due to difference in results of MS-MSD. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples C0AC9, C0AC9DL, C0O14, C0O14DL, C0O15, C0O15DL, C0O16, C0O16DL, C0O17, C0O17DL, C0O18, C0O18DL, C0O19, C0O20, C0O21, C0O22, C0O23, C0O24 and C0O24DL were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date