| CHEMTECH PROJECT NUMBER: Q1626 | MATRIX: Solid |
| --- | --- |
| METHOD: 8081B/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1626-01MS} with File ID: PL094837.D recoveries met the requirements for all compounds except for 4,4-DDD[300%], 4,4-DDE[139%], 4,4-DDT[307%], Dieldrin[239%], Endosulfan sulfate[141%], Endrin[171%], Heptachlor epoxide[169%] due to matrix interference. The MSD {Q1626-01MSD} with File ID: PL094838.D recoveries met the acceptable requirements except for 4,4-DDD[287%], 4,4-DDT[302%], Dieldrin[233%], Endrin[166%], Heptachlor epoxide[161%] due to matrix interference . The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples CO-32-1 was diluted due to bad matrix.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date