| CHEMTECH PROJECT NUMBER: Q1664 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BG030525.M) for Benzoic acid, 2,6-Dinitrotoluene, 2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenol, Butylbenzlphthalate, these compounds are passing on Linear Regression and 2-Nitrophenol, 2-Nitroaniline, are passing on Quadratic regression. The Continuous Calibration File ID BG064130.D met the requirements except for Hexachlorocyclopentadiene, The associate samples have no positive hit for this compound therefore no corrective action was taken. . The Continuous Calibration File ID BG064164.D met the requirements except for 2,4-Dinitrophenol, 2-Nitrophenol, 4,6-Dinitro-2-methylphenol and Hexachlorocyclopentadiene, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for P001-BBDGA-005-01 [2,4 and6-Tribromophenol - 138%], as per method two surrogates are allowed to failed, therefore no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1664-05MS} with File ID: BG064135.D recoveries met the requirements for all compounds except for 1,4-Dioxane[30%], 2,3,4,6-Tetrachlorophenol[122%], 2,4,5-Trichlorophenol[123%], 2,4,6-Trichlorophenol[123%] and Hexachlorocyclopentadiene[230%], due to matrix interference, therefore no corrective action is required. The MSD {Q1664-06MSD} with File ID: BG064136.D recoveries met the acceptable requirements except for 1,2,4,5-Tetrachlorobenzene[103%], 1,4-Dioxane[32%], 2,3,4,6-Tetrachlorophenol[129%], 2,4,5-Trichlorophenol[125%], 2,4,6-Trichlorophenol[127%], Atrazine[164%], Hexachlorocyclopentadiene[240%] and Nitrobenzene[114%], due to matrix interference, therefore no corrective action is required. The Blank Spike for {PB167393BS} with File ID: BG064173.D met requirements for all samples except for 2,4,5-Trichlorophenol[107%], 2-Nitrophenol[118%], 3+4-Methylphenols[109%], 4,6-Dinitro-2-methylphenol[133%], Atrazine[135%], Butylbenzylphthalate[108%] and Hexachlorocyclopentadiene[230%], The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date