| CHEMTECH PROJECT NUMBER: Q1664 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF031025.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The % RSD is greater than 20% in the Initial Calibration (8270-BM031325.M) for 4-Nitroaniline, this compound is passing on Linear Regression, and 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol passing on Quadratic regression. The Continuous Calibration File ID BF142175.D met the requirements except for Benzaldehyde, The associate samples have no positive hit for these compounds therefore no corrective action was taken. The Continuous Calibration File ID BM049784.D met the requirements except for Terphenyl-d14, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for PB167369BS [Terphenyl-d14 - 112%], marginally out of QC limits, therefore no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1664-02MS} with File ID: BM049797.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[111%], 2-Chlorophenol[111%], 3+4-Methylphenols[117%], Atrazine[172%] and Benzaldehyde[111%], due to matrix interference therefore no corrective action is required . The MSD {Q1664-03MSD} with File ID: BM049798.D recoveries met the acceptable requirements except for 2,2-oxybis(1-Chloropropane)[117%], 2,3,4,6-Tetrachlorophenol[117%], 2-Chlorophenol[111%], 3+4-Methylphenols[122%], 4-Chlorophenyl-phenylether[111%], Atrazine[178%] and Benzaldehyde[117%], due to matrix interference, therefore no corrective action is required. The Blank Spike for {PB167369BS} with File ID: BF142177.D met requirements for all samples except for 2,4,5-Trichlorophenol[106%], 2,4-Dinitrotoluene[112%], 2,6-Dinitrotoluene[106%], 2-Chloronaphthalene[100%], 3,3-Dichlorobenzidine[19%], 4,6-Dinitro-2-methylphenol[118%], 4-Chlorophenyl-phenylether[106%], Acenaphthylene [106%], Anthracene[106%], Benzo(a)anthracene[106%], Benzo(a)pyrene[112%], Carbazole[100%], Chrysene[106%], Dimethylphthalate[100%], Hexachlorobenzene [106%], Hexachlorobutadiene[106%], Hexachlorocyclopentadiene[185%], Isophorone[100%], N-Nitrosodiphenylamine[100%], Pentachlorophenol[106%] and Pyrene[106%], The associate samples have no positive hit for these compounds therefore no corrective action was taken. The RPD for {Q1664-03MSD} with File ID: BM049798.D met criteria except for 4-Chloroaniline[22%],due to matrix interference, no corrective action was taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for P001-BBDGA-001-02, P001-BBDGA-002-01, P001-BBDGA-005-01 and P001-BBDGA-006-01, The sample was reanalyzed to confirm internal standard failure, both run were reported in Hard Copy. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date