| CHEMTECH PROJECT NUMBER: Q1665 | MATRIX: Air |
| --- | --- |
| METHOD: TO-15 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID VL042326.D met the requirements except for Methylene Chloride and tert-Butyl alcohol are failing low |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike Duplicate met requirements for all samples .The Blank Spike for {VL0411ABS01} with File ID: VL042314.D met requirements for all samples except for 1,1,2,2-Tetrachloroethane[132%], 1,2-Dichlorobenzene[132%],these compounds did not meet the NJDKQP criteria and the in-house criteria are failing high but no positive hit in associate samples therefore no corrective action taken. The Blank Spike for {VL0415ABS01} with File ID: VL042328.D met requirements for all samples except for Methylene Chloride[67%], tert-Butyl Alcohol[68%],these compounds did not meet the NJDKQP criteria and the in-house criteria are failing low but no positive hit in associate samples therefore no corrective action taken. The RPD for {Q1665-01DUP} with File ID: VL042291.D met criteria except for m/p-Xylene[200%],due to difference in results of original and DUP. The RPD for {Q1665-03DUP} with File ID: VL042318.D met criteria except for m/p-Xylene[200%],due to difference in results of original and DUP. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: The Holding Times were met for all analysis. |  |  |  |

ADDITIONAL COMMENTS:

Samples IA-B10-2-3242025, IA-B10-4-3242025, IA-B10-4-3242025DL, IA-B10-1-3242025, IA-B14-1-3242025 and IA-B10-3-3242025 were diluted due to high concentrations.

The Manual Integrations are performed for the followings.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The Sample #AA-DUP-1-3242025,IA-B10-2-3242025,AA-1-3242025,AA-1-3242025,IA-B9-2-3242025,IA-B10-4-3242025,IA-B14-1-3242025,IA-B14-1-3242025DL,IA-B10-3-3242025 have the concentration of target compound belowMethod detection limits, therefore it is not reported as Hit in Form1.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date