| CHEMTECH PROJECT NUMBER: Q1690 | MATRIX: Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BG030525.M) for 2,6-Dinitrotoluene, 2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenol, Butylbenzlphthalate, these compounds are passing on Linear Regression and 2-Nitrophenol, 2-Nitroaniline, are passing on Quadratic regression. The Continuous Calibration File ID BG064164.D met the requirements except for 2,4-Dinitrophenol,2-Nitrophenol,4,6-Dinitro-2-methylphenol and Hexachlorocyclopentadiene, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for PB167451BL [2,4,6-Tribromophenol - 112%], PB167451BS [2,4,6-Tribromophenol - 125%], PB167451BSD [2,4,6-Tribromophenol - 124%], this compound did not meet the NJDKQP criteria but met the in-house criteria, while MW112010 [2,4,6-Tribromophenol - 0%, 2-Fluorophenol - 0% and Phenol-d6 - 0%], these compounds did not meet the NJDKQP criteria and in-house criteria, but matrix interference Due to the limited volume availability of this sample it can not be re-extracted and re-analyzing the sample of surrogate failure therefore no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB167451BS} with File ID: BG064175.D met requirements for all samples except for 3,3-Dichlorobenzidine[54%], 3-Nitroaniline[44%], and 4-Chloroaniline[25%],these compounds did not meet the NJDKQP criteria but met the in-house criteria, while Hexachlorocyclopentadiene[250%], 2,4-Dimethylphenol[165%], 4,6-Dinitro-2-methylphenol[137%], Atrazine[143%], these compounds did not meet the NJDKQP criteria and in-house criteria, The associate samples have no positive hit for these compounds therefore no corrective action was taken. The Blank Spike Duplicate met requirements for all samples . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date