| CHEMTECH PROJECT NUMBER: Q1711 | MATRIX: Water |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1711-02MS} with File ID: VX045591.D recoveries met the requirements for all compounds except for cis-1,2-Dichloroethene[30%] this compound did not meet the NJDKQP criteria and in-house criteria, due to matrix interference. The MSD {Q1711-03MSD} with File ID: VX045577.D recoveries met the acceptable requirements except for cis-1,2-Dichloroethene[50%] this compound did not meet the NJDKQP criteria and in-house criteria, due to matrix interference. The RPD for {Q1711-03MSD} with File ID: VX045577.D met criteria except for cis-1,2-Dichloroethene[50%] this compound did not meet the NJDKQP criteria and in-house criteria due, to difference in results of MS and MSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW-18B-56-040225, MW-17B-55-040225 was diluted due to high concentrations of compound Trichloroethene.

Trip Blank was not provided with this set of samples.

For Sample #01 & 04 at the time of fax Sequence Processed with wrong method after further review it is corrected in Hardcopy therefore fax and Hardcopy data will not match.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date