CASE NARRATIVE

### Woodard & Curran

### Project Name: US EPA Riverside Industrial Park Superfund Site

### Project # N/A

### Chemtech Project # Q1726

Test Name: TO-15

A. Number of Samples and Date of Receipt:

7 Air samples were received on 04/03/2025.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

TO-15. This data package contains results for TO-15.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_L were done using GC column RTX-1, which is 60 meters, 0.32 mm id, 1.0 um df, Restek Cat. #10157. The Trap was supplied by Entech, glass bead and Tenax , Entech 7100A Preconcentrator.The analysis of TO-15 was based on method TO-15.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria.

The Internal Standards Areas met the acceptable requirements.

The Retention Times were acceptable for all samples.

The RPD for {Q1726-06DL2DUP} with File ID: VL042393.D met criteria except for Ethanol[27.5%] due to difference in results of Original and DUP.

The Blank Spike for {VL0417ABS01} with File ID: VL042358.D met requirements for all samples except for Ethanol[32%] this compound did not meet the NJDKQP criteria and the in-house criteria.

The Blank Spike for {VL0418ABS01} with File ID: VL042385.D met requirements for all samples except for Ethanol[23%], Heptane[68%], Methyl methacrylate[69%] and Tetrahydrofuran[69%],these compounds did not meet the NJDKQP criteria and the in-house criteria also it is failing low but only dilution samples analyzed under this Blank spike therefore no corrective action taken.

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration with dated 04/17/2025 with L Instrument met the requirements except for Ethanol.

The Continuous Calibration File ID VL042383.D met the requirements except for Ethanol, 1,4-Dioxane,4-Methyl-2-Pentanone and Tetrahydrofuran are failing low but only dilution samples analyzed under this Continuous Calibration therefore no corrective action taken.

The Tuning criteria met requirements.

Samples IA-DUP-3-03312025, IA-DUP-3-03312025DL, IA-B2-6-03312025, IA-B2-6-03312025DL, IA-B2-2-03312025, IA-B2-2-03312025DL, IA-B2-5-03312025, IA-B2-5-03312025DL, IA-B2-1-03312025, IA-B2-1-03312025DL, IA-B2-3-03312025, IA-B2-3-03312025DL, IA-B2-4-03312025 and IA-B2-4-03312025DL were diluted due to high concentrations.

E. Additional Comments:

This data Package has been revised to include Parameter Bromoethene, Ethanol and

isopropyl alcohol.

This data Package has been revised to include DUP raw data.

The not QT review data is reported in the Miscellaneous.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The Sample #  IA-DUP-3-03312025DL, IA-DUP-3-03312025DL2, IA-B2-6-03312025DL, IA-B2-2-03312025DL, IA-B2-1-03312025DL2 and IA-B2-4-03312025DL2 have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1.

The Manual Integrations are performed for the followings.

| Manual Integration Report |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sequence | VL041725 | Instrument | MSVOA_l |  |  |  |  |

| Sample ID | File ID | Parameter | Review By | Review On | Supervised By | Supervised On | Reason |
| --- | --- | --- | --- | --- | --- | --- | --- |
| VSTDICCC010 | VL042349.D | m/p-Xylene | SAM | 4/18/2025 4:19:07 PM | MMDadoda | 4/18/2025 10:45:20 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | 1,4-Dioxane | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | cis-1,3-Dichloropropene | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | Ethanol | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | Heptane | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | m/p-Xylene | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042350.D | Trichloroethene | SAM | 4/18/2025 4:19:12 PM | MMDadoda | 4/18/2025 10:45:17 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | 1,1,2-Trichloroethane | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | 1,3-Butadiene | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | 1,4-Dioxane | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | 2,2,4-Trimethylpentane | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | cis-1,3-Dichloropropene | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | Cyclohexane | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | Ethanol | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | Heptane | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | m/p-Xylene | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | Methyl Methacrylate | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042351.D | t-1,3-Dichloropropene | SAM | 4/18/2025 4:19:17 PM | MMDadoda | 4/18/2025 10:45:15 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | 1,1,2-Trichloroethane | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | 1,2-Dichloropropane | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | 1,4-Dioxane | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | 4-Methyl-2-Pentanone | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | Chlorobenzene | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | cis-1,3-Dichloropropene | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | Ethanol | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | Heptane | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | m/p-Xylene | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | Methyl Methacrylate | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | t-1,3-Dichloropropene | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042352.D | Tetrahydrofuran | SAM | 4/18/2025 4:20:33 PM | MMDadoda | 4/18/2025 10:45:07 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042353.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:20:37 PM | MMDadoda | 4/18/2025 10:45:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042353.D | 1,1,2,2-Tetrachloroethane | SAM | 4/18/2025 4:20:37 PM | MMDadoda | 4/18/2025 10:45:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042353.D | 1,2-Dibromoethane | SAM | 4/18/2025 4:20:37 PM | MMDadoda | 4/18/2025 10:45:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042353.D | Tetrachloroethene | SAM | 4/18/2025 4:20:37 PM | MMDadoda | 4/18/2025 10:45:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042353.D | Trichloroethene | SAM | 4/18/2025 4:20:37 PM | MMDadoda | 4/18/2025 10:45:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042354.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:19:24 PM | MMDadoda | 4/18/2025 10:45:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042354.D | 1,1,2,2-Tetrachloroethane | SAM | 4/18/2025 4:19:24 PM | MMDadoda | 4/18/2025 10:45:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042354.D | Carbon Tetrachloride | SAM | 4/18/2025 4:19:24 PM | MMDadoda | 4/18/2025 10:45:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042354.D | Tetrachloroethene | SAM | 4/18/2025 4:19:24 PM | MMDadoda | 4/18/2025 10:45:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042354.D | Trichloroethene | SAM | 4/18/2025 4:19:24 PM | MMDadoda | 4/18/2025 10:45:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC015 | VL042355.D | m/p-Xylene | SAM | 4/18/2025 4:21:35 PM | MMDadoda | 4/18/2025 10:44:59 PM | Peak Integrated by Software incorrectly |
| VSTDICV010 | VL042356.D | m/p-Xylene | SAM | 4/18/2025 4:20:42 PM | MMDadoda | 4/18/2025 10:44:57 PM | Peak Integrated by Software incorrectly |
| VL0417ABS01 | VL042358.D | m/p-Xylene | SAM | 4/18/2025 4:21:39 PM | MMDadoda | 4/18/2025 10:44:55 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | 1,2-Dichloroethane | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | Carbon Tetrachloride | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | Chlorodifluoromethane | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | n-Butylbenzene | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | Naphthalene,2-methyl- | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | Propene | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-01 | VL042366.D | Tetrachloroethene | SAM | 4/18/2025 4:19:35 PM | MMDadoda | 4/18/2025 10:44:26 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Carbon Tetrachloride | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Chlorodifluoromethane | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Chloroform | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Cyclohexane | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Propene | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-02 | VL042367.D | Tetrachloroethene | SAM | 4/18/2025 4:20:08 PM | MMDadoda | 4/18/2025 10:45:51 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Carbon Tetrachloride | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Chlorodifluoromethane | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Chloroform | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | n-Butylbenzene | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Naphthalene,2-methyl- | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Propene | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | tert-Butyl alcohol | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-03 | VL042368.D | Tetrachloroethene | SAM | 4/18/2025 4:21:50 PM | MMDadoda | 4/18/2025 10:44:24 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | Carbon Tetrachloride | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | Chlorodifluoromethane | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | Propene | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | tert-Butyl alcohol | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-04 | VL042369.D | Tetrachloroethene | SAM | 4/18/2025 4:21:55 PM | MMDadoda | 4/18/2025 10:44:18 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | 1,1,1-Trichloroethane | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | Chlorodifluoromethane | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | Cyclohexane | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | Propene | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | tert-Butyl alcohol | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-05 | VL042370.D | Tetrachloroethene | SAM | 4/18/2025 4:21:04 PM | MMDadoda | 4/18/2025 10:44:08 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | Chlorodifluoromethane | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | n-Butylbenzene | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | Propene | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | tert-Butyl alcohol | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | Tetrachloroethene | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-06 | VL042371.D | Trichloroethene | SAM | 4/18/2025 4:20:03 PM | MMDadoda | 4/18/2025 10:44:05 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | 4-Methyl-2-Pentanone | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | Carbon Tetrachloride | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | Chlorodifluoromethane | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | Chloroform | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | n-Butylbenzene | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | Propene | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | tert-Butyl alcohol | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07 | VL042372.D | Tetrachloroethene | SAM | 4/18/2025 4:20:49 PM | MMDadoda | 4/18/2025 10:44:02 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL | VL042373.D | 2,2,4-Trimethylpentane | SAM | 4/18/2025 4:22:12 PM | MMDadoda | 4/18/2025 10:43:58 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL | VL042373.D | Heptane | SAM | 4/18/2025 4:22:12 PM | MMDadoda | 4/18/2025 10:43:58 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL | VL042373.D | m/p-Xylene | SAM | 4/18/2025 4:22:12 PM | MMDadoda | 4/18/2025 10:43:58 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL | VL042373.D | Propene | SAM | 4/18/2025 4:22:12 PM | MMDadoda | 4/18/2025 10:43:58 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL | VL042376.D | Benzene | SAM | 4/18/2025 4:19:57 PM | MMDadoda | 4/18/2025 10:43:47 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL | VL042376.D | Cyclohexane | SAM | 4/18/2025 4:19:57 PM | MMDadoda | 4/18/2025 10:43:47 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL | VL042376.D | Propene | SAM | 4/18/2025 4:19:57 PM | MMDadoda | 4/18/2025 10:43:47 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | Benzene | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | Cyclohexane | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | Heptane | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | m/p-Xylene | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | n-propylbenzene | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL | VL042377.D | Propene | SAM | 4/18/2025 4:21:15 PM | MMDadoda | 4/18/2025 10:43:43 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL | VL042378.D | 1,3,5-Trimethylbenzene | SAM | 4/18/2025 4:21:23 PM | MMDadoda | 4/18/2025 10:43:37 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL | VL042378.D | 2,2,4-Trimethylpentane | SAM | 4/18/2025 4:21:23 PM | MMDadoda | 4/18/2025 10:43:37 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL | VL042378.D | Heptane | SAM | 4/18/2025 4:21:23 PM | MMDadoda | 4/18/2025 10:43:37 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL | VL042378.D | m/p-Xylene | SAM | 4/18/2025 4:21:23 PM | MMDadoda | 4/18/2025 10:43:37 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL | VL042378.D | Propene | SAM | 4/18/2025 4:21:23 PM | MMDadoda | 4/18/2025 10:43:37 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL | VL042379.D | 2,2,4-Trimethylpentane | SAM | 4/18/2025 4:19:52 PM | MMDadoda | 4/18/2025 10:43:34 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL | VL042379.D | Benzene | SAM | 4/18/2025 4:19:52 PM | MMDadoda | 4/18/2025 10:43:34 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL | VL042379.D | Heptane | SAM | 4/18/2025 4:19:52 PM | MMDadoda | 4/18/2025 10:43:34 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL | VL042379.D | n-propylbenzene | SAM | 4/18/2025 4:19:52 PM | MMDadoda | 4/18/2025 10:43:34 PM | Peak Integrated by Software incorrectly |
| Q1726-05DL | VL042380.D | Heptane | SAM | 4/18/2025 4:19:46 PM | MMDadoda | 4/18/2025 10:43:30 PM | Peak Integrated by Software incorrectly |
| Q1726-05DL | VL042380.D | Propene | SAM | 4/18/2025 4:19:46 PM | MMDadoda | 4/18/2025 10:43:30 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | 2-Butanone | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | Heptane | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | Hexane | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | m/p-Xylene | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | Propene | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL | VL042381.D | Toluene | SAM | 4/18/2025 4:19:42 PM | MMDadoda | 4/18/2025 10:45:53 PM | Peak Integrated by Software incorrectly |

| Manual Integration Report |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sequence | VL041825 | Instrument | MSVOA_l |  |  |  |  |

| Sample ID | File ID | Parameter | Review By | Review On | Supervised By | Supervised On | Reason |
| --- | --- | --- | --- | --- | --- | --- | --- |
| VSTDCCC010 | VL042383.D | m/p-Xylene | SAM | 4/21/2025 7:55:57 AM | MMDadoda | 4/21/2025 3:26:35 PM | Peak Integrated by Software incorrectly |
| VL0418ABS01 | VL042385.D | m/p-Xylene | SAM | 4/21/2025 7:56:01 AM | MMDadoda | 4/21/2025 3:26:35 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL2 | VL042386.D | m/p-Xylene | SAM | 4/21/2025 7:56:07 AM | MMDadoda | 4/21/2025 3:26:35 PM | Peak Integrated by Software incorrectly |
| Q1726-07DL2 | VL042386.D | Toluene | SAM | 4/21/2025 7:56:07 AM | MMDadoda | 4/21/2025 3:26:35 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL2 | VL042387.D | 1,3,5-Trimethylbenzene | SAM | 4/21/2025 7:56:30 AM | MMDadoda | 4/21/2025 3:26:37 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL2 | VL042387.D | Ethyl Benzene | SAM | 4/21/2025 7:56:30 AM | MMDadoda | 4/21/2025 3:26:37 PM | Peak Integrated by Software incorrectly |
| Q1726-01DL2 | VL042387.D | Heptane | SAM | 4/21/2025 7:56:30 AM | MMDadoda | 4/21/2025 3:26:37 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL2 | VL042388.D | 2,2,4-Trimethylpentane | SAM | 4/21/2025 7:56:12 AM | MMDadoda | 4/21/2025 3:26:38 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL2 | VL042388.D | Ethyl Benzene | SAM | 4/21/2025 7:56:12 AM | MMDadoda | 4/21/2025 3:26:38 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL2 | VL042388.D | Heptane | SAM | 4/21/2025 7:56:12 AM | MMDadoda | 4/21/2025 3:26:38 PM | Peak Integrated by Software incorrectly |
| Q1726-02DL2 | VL042388.D | o-Xylene | SAM | 4/21/2025 7:56:12 AM | MMDadoda | 4/21/2025 3:26:38 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL2 | VL042389.D | 2,2,4-Trimethylpentane | SAM | 4/21/2025 7:56:35 AM | MMDadoda | 4/21/2025 3:26:40 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL2 | VL042389.D | Ethyl Benzene | SAM | 4/21/2025 7:56:35 AM | MMDadoda | 4/21/2025 3:26:40 PM | Peak Integrated by Software incorrectly |
| Q1726-04DL2 | VL042389.D | Heptane | SAM | 4/21/2025 7:56:35 AM | MMDadoda | 4/21/2025 3:26:40 PM | Peak Integrated by Software incorrectly |
| Q1726-03DL2 | VL042390.D | Toluene | SAM | 4/21/2025 7:56:40 AM | MMDadoda | 4/21/2025 3:26:42 PM | Peak Integrated by Software incorrectly |
| Q1726-05DL2 | VL042391.D | m/p-Xylene | SAM | 4/21/2025 7:56:45 AM | MMDadoda | 4/21/2025 3:26:44 PM | Peak Integrated by Software incorrectly |
| Q1726-05DL2 | VL042391.D | Toluene | SAM | 4/21/2025 7:56:45 AM | MMDadoda | 4/21/2025 3:26:44 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL2 | VL042392.D | Toluene | SAM | 4/21/2025 7:56:17 AM | MMDadoda | 4/21/2025 3:26:46 PM | Peak Integrated by Software incorrectly |
| Q1726-06DL2DUP | VL042393.D | Toluene | SAM | 4/21/2025 7:56:21 AM | MMDadoda | 4/21/2025 3:26:48 PM | Peak Integrated by Software incorrectly |

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________