CASE NARRATIVE

### Ardmore Chemical

### Project Name: PVSC Monthly 2025

### Project # N/A

### Chemtech Project # Q1833

Test Name: VOC-PP

A. Number of Samples and Date of Receipt:

2 Water samples were received on 04/17/2025.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

BOD5, Cyanide, Mercury, Metals Group2, Metals Group3, Metals ICP-Group, SVOCMS Group1, TSS and VOC-PP. This data package contains results for VOC-PP.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-PP was based on method 624.1.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria.

The Internal Standards Areas met the acceptable requirements.

The Retention Times were acceptable for all samples.

The RPD met criteria.

The Blank Spike met requirements for all samples.

The Blank Spike Duplicate met requirements for all samples.

The Blank analysis did not indicate the presence of lab contamination.

The %RSD is greater than 35% in the Initial Calibration method (624N042325W.M) for Acrolein is passing on Linear Regression.

The Tuning criteria met requirements.

Sample EFF-WASTE-WATER was diluted due foamy nature of the sample.

E. Additional Comments:

As per method, MS/MSD is required to be performed with the sample analysis. However, Lab did not receive sufficient volume to perform the MS/MSD therefore MS/MSD were not performed for this project. However, Lab has performed LCS/LCSD instead.

Sample # EFF-WASTE-WATER initially analyzed in sequence VN042125  where Method was Failing as a corrective action sample reanalyzed in sequence VN042325 under passing method and reported in hardcopy, therefore fax and hardcopy data will not match.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <35% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 35% for the Initial Calibration curve for SW-846 analysis.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________