| ORDER ID: Q1872 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. The Initial Calibration met the requirements . The Continuous Calibration met the requirements . |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1901-04MS} with File ID: FC068804.D recoveries met the requirements for all compounds except for [n-Tetracosane (C24)- 39%] due to matrix interference and [Naphthalene (C11.7)- 1%, 2-methylnaphthalene (C12.89)- 0%] these analyzes compounds are only being monitoring in aliphatic fraction. The MSD {Q1901-04MSD} with File ID: FC068805.D recoveries met the requirements for all compounds except for [n-Tetracosane (C24)- 39%] due to matrix interference and [Naphthalene (C11.7)- 1%, 2-methylnaphthalene (C12.89)- 0%] these analyzes compounds are only being monitoring in aliphatic fraction. The MS {Q1901-04MS} with File ID: FD049330.D recoveries met the requirements for all compounds except for [Chrysene (C27.41)-196%, Bnezo[k]fluoranthene (C30.14)-183%, Dibenz[a,h]anthracene (C30.36)-200%, benzo[b]fluoranthene (C30.41)-180%, Indeno[1,2,3-cd]pyrene (C35.01)-187%] due to matrix interference. The MSD {Q1901-04MSD} with File ID: FD049331.D recoveries met the requirements for all compounds except for [Benzo[a]anthracene (C26.37)-199%, Chrysene (C27.41)-195%, Bnezo[k]fluoranthene (C30.14)-183%, Dibenz[a,h]anthracene (C30.36)-200%, benzo[b]fluoranthene (C30.41)-180%, Indeno[1,2,3-cd]pyrene (C35.01)-188%] due to matrix interference. The RPD for {Q1901-04MSD} with File ID: FD049331.D met criteria except for Benzo[a]anthracene (C26.37)- 68.92% due to difference in results of MS and MSD. The Blank Spike {PB167903BS} with File ID: FC068793.D met requirements for all samples except for [Naphthalene (C11.7)-0%, 2-methylnaphthalene (C12.89)-0% ] these analyzes compounds are only being monitoring in aliphatic friction. The Blank Spike Duplicate{PB167903BSD} with File ID: FC068794.D met requirements for all samples except for [Naphthalene (C11.7)-0%, 2-methylnaphthalene (C12.89)-0% ] these analyzes compounds are only being monitoring in aliphatic friction. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples HW0425-PT-NJEPH-SOIL and HW0425-PT-NJEPH-SOILDL were diluted due to high concentrations for Aliphatic compounds.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date