| CHEMTECH PROJECT NUMBER: Q1901 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration (Method 8270-BF043025.M) for Di-n-octyl phthalate, Bis(2-ethylhexyl)phthalate, Butylbenzylphthalate, 4,6-Dinitro-2-methylphenol, 4-Nitroaniline, 2,4-Dinitrotoluene, 3-Nitroaniline, 2,6-Dinitrotoluene, 2-Nitroaniline, these Compounds are passing on Linear regression, and 2,4-Dinitrophenol, 2-Nitrophenol, these compound are passing on Quadratic regression. The % RSD is greater than 20% in the Initial Calibration (8270-BM042825.M) for 2,4-Dinitrophenol and 4-Nitrophenol these compound are passing on Linear Regression. The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1901-02MS} with File ID: BF142269.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[65%], 3-Nitroaniline[50%], 4-Chloroaniline[21%], Benzo(g,h,i)perylene[65%] and Pyrene[64%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The MSD {Q1901-02MSD} with File ID: BF142270.D recoveries met the acceptable requirements except for 3-Nitroaniline[50%], 4-Chloroaniline[24%], Benzo(g,h,i)perylene[60%], Dibenz(a,h)anthracene[65%], Indeno(1,2,3-cd)pyrene[65%] and Pyrene[64%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The Blank Spike for {PB167798BS} with File ID: BF142252.D met requirements for all samples except for 3-Nitroaniline[39%],4-Chloroaniline[13%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while 3,3-Dichlorobenzidine[23%], this compound did not meet the NJDKQP criteria and in-house criteria but marginally biased low therefore no corrective action was taken. The Blank Spike Duplicate for {PB167798BSD} with File ID: BF142253.D met requirements for all samples except for 3-Nitroaniline[45%] and 4-Chloroaniline[18%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while 3,3-Dichlorobenzidine[28%], this compound did not meet the NJDKQP criteria and in-house criteria but marginally biased low therefore no corrective action was taken. The Blank Spike for {PB167779BS} with File ID: BM050064.D met requirements for all samples except for 3,3-Dichlorobenzidine[49%], 3-Nitroaniline[47%] and 4-Chloroaniline[36%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The RPD for {PB167798BSD} with File ID: BF142253.D met criteria except for 4-Chloroaniline[31%], due to difference in results of BS and BSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples B-170-SB01 analyzed with direct 5x dilution due to dirty, concentrated and viscous matrix.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date