| CHEMTECH PROJECT NUMBER: Q1907 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BM042825.M) for 2,4-Dinitrophenol and 4-Nitrophenol, these compound are passing on Linear Regression. The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1914-10MS} with File ID: BM050083.D recoveries met the requirements for all compounds except for 1,1-Biphenyl[51%], 2,2-oxybis(1-Chloropropane)[61%], 2,4-Dichlorophenol[128%], 2-Methylnaphthalene[-6%], 2-Nitroaniline[67%], 3,3-Dichlorobenzidine[12%], 4-Nitroaniline[54%], Acenaphthene[67%], Acenaphthylene[72%], Acetophenone[333%], Caprolactam[161%], Dimethylphthalate[67%], Di-n-butylphthalate[64%], Hexachlorobutadiene[128%], Hexachloroethane[283%], Isophorone[156%], Naphthalene[67%] and Nitrobenzene[128%] due to matrix interference. The MSD {Q1914-11MSD} with File ID: BM050084.D recoveries met the acceptable requirements except for 1,1-Biphenyl[56%], 2,2-oxybis(1-Chloropropane)[61%], 2,4-Dichlorophenol[139%], 2-Methylnaphthalene[6%], 3,3-Dichlorobenzidine[12%], Acenaphthylene[78%], Acetophenone[350%], bis(2-Ethylhexyl)phthalate[41%], Caprolactam[156%], Di-n-butylphthalate[64%], Hexachlorobutadiene[139%], Hexachloroethane[289%], Isophorone[161%] and Nitrobenzene[139%] due to matrix interference. The Blank Spike for {PB167803BS} with File ID: BM050081.D met requirements for all samples except for 3,3-Dichlorobenzidine[39%] but associated samples have not positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples CO-8R-WC was diluted due to concentrated and viscous matrix.

Sample CO-8R-WC was diluted due to high concentration.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date