| ORDER ID: Q2008 | MATRIX: Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF050525.M) for 2-Nitrophenol, 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol, Butylbenzlphthalate, Bis(2-ethylhexyl)phthalate, Di-n-octyl phthalate, these compound are passing on Linear Regression The Continuous Calibration File ID BF142336.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4,5-Trichlorophenol,2,4,6-Trichlorophenol,2,4-Dichlorophenol,2,4-Dinitrotoluene,2,6-Dinitrotoluene,2-Nitroaniline,3,3-Dichlorobenzidine,3-Nitroaniline,4-Nitroaniline,4-Nitrophenol,Atrazine,Caprolactam,Di-n-butylphthalate,Hexachlorocyclopentadiene,Nitrobenzene,Pentachlorophenol,Phenol,2,4,6-Tribromophenol and Nitrobenzene-d5 . But associated samples have not positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for IDW-AQ-DRUM-633-05092025 [2-Fluorophenol - 13%- Phenol-d6 - 9%]. 2-Fluorophenol did not meet the NJDKQP criteria but met the in-house criteria. AS per method one Acid surrogate is allowed to fail, Therefore no corrective action required. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1993-02MS} with File ID: BF142341.D recoveries met the requirements for all compounds except for 2-Methylphenol[64%], 3-Nitroaniline[50%], 4-Chloroaniline[29%], Benzo(g,h,i)perylene[64%], Caprolactam[19%] and Diethylphthalate[60%] . These compounds did not meet the NJDKQP criteria but met the in-house criteria, due to matrix interference. The MSD {Q1993-03MSD} with File ID: BF142342.D recoveries met the acceptable requirements except for 2-Methylphenol[68%], 3-Nitroaniline[52%], 4-Chloroaniline[30%], Benzo(g,h,i)perylene[66%] and Diethylphthalate[68%]. These compounds did not meet the NJDKQP criteria but met the in-house criteria, due to matrix interference. The Blank Spike for {PB167951BS} with File ID: BF142338.D met requirements for all samples except for 4-Chloroaniline[25%] . This compound did not meet the NJDKQP criteria but met the in-house criteria. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date