| ORDER ID: Q2032 | MATRIX: Solid, Water |
| --- | --- |
| METHOD: 6010D,7470A,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution met criteria for all samples. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (3859MS) analysis met criteria for all samples except for Mercury due to matrix interference. The Matrix Spike (TP-24MS) analysis met criteria for all samples except for Antimony, Barium, Magnesium, Potassium, Selenium, Silver, Sodium, Vanadium due to matrix interference. The Matrix Spike (TW-WTS-08MS) analysis met criteria for all samples except for Iron, Silver due to matrix interference. The Matrix Spike Duplicate (3859MSD) analysis met criteria for all samples except for Mercury due to matrix interference. The Matrix Spike Duplicate (TP-24MSD) analysis met criteria for all samples except for Antimony, Magnesium, Manganese, Potassium, Vanadium, Zinc due to matrix interference. The Matrix Spike Duplicate (TW-WTS-08MSD) analysis met criteria for all samples except for Silver due to matrix interference. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (3859MSD) analysis met criteria for all samples except for Mercury due to matrix interference. The Duplicate (TP-24DUP) analysis met criteria for all samples except for Lead due to matrix interference. The Duplicate (TP-24MSD) analysis met criteria for all samples except for Aluminum, Calcium, Iron, Magnesium, Manganese, Sodium, Vanadium due to matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

_______________________________________				__________________________

QA REVIEW								Date