| ORDER ID: Q2032 | MATRIX: /Water |
| --- | --- |
| METHOD: 8260-Low,8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID VY022329.D met the requirements except for Chloroethane and Vinyl Chloride . But associated samples have not positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for TP-11 [Toluene-d8 - 125%]. Due to high concentration of compounds, this sample required MEOH. Therefore, sample was reanalyzed in MEOH and reported. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MSD {Q2032-06MSD} with File ID: VY022349.D recoveries met the acceptable requirements except for 1,1,2-Trichlorotrifluoroethane[148%], Acetone[185%], Methyl Acetate[222%] and Methylcyclohexane[150%] due to matrix interference. The Blank Spike for {VY0520SBS01} with File ID: VY022331.D met requirements for all samples except for Bromomethane[190%], Chloroethane[175%], Chloromethane[149%], Trichlorofluoromethane[143%] and Vinyl chloride[153%]. But associated samples have not positive hit for these compounds therefore no corrective action was taken. The Blank Spike Duplicate met requirements for all samples . The RPD for {Q2032-06MSD} with File ID: VY022349.D met criteria except for 1,1,2,2-Tetrachloroethane[35%], 1,1,2-Trichlorotrifluoroethane[29%], 1,2,3-Trichlorobenzene[63%], 1,2,4-Trichlorobenzene[50%], Bromochloromethane[21%], Bromomethane[40%], Carbon Tetrachloride[36%], Chloroform[24%], cis-1,2-Dichloroethene[21%], cis-1,3-Dichloropropene[25%], Cyclohexane[25%], Dibromochloromethane[21%], Isopropylbenzene[46%], Methyl Acetate[27%], Methyl tert-butyl Ether[27%], Methylcyclohexane[40%], Methylene Chloride[27%], t-1,3-Dichloropropene[23%] and Tetrachloroethene[36%] due to difference in MS and MSD concentrations. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for TP-24MSD. |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Sample TP-11 was diluted due to high concentration.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date