| ORDER ID: Q2161 | MATRIX: Solid |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID VX046433.D met the requirements except for Bromoform,cis-1,3-Dichloropropene,Dibromochloromethane,Methyl Acetate and t-1,3-Dichloropropene, failing high but no positive hit in associated samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for B27-SOIL-SAMPLE [Dibromofluoromethane - 42%], as VIAL A analyzed but did not purge As a corrective action VIAL B analyzed but surrogate Fail, therefore VIAL B reported as final analysis, and B28-SOIL-SAMPLE [1,2-Dichloroethane-d4 - 54%, Dibromofluoromethane - 42%], VIAL B analyzed but did not purge As a corrective action VIAL A analyzed but Surrogate Fail, therefore VIAL A reported as final analysis. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {VY0530SBS01} with File ID: VY022474.D met requirements for all samples except for Bromomethane[144%], Chloroethane[136%], failing high but no positive hit in associated samples therefore no corrective action taken. The Blank Spike Duplicate for {VY0530SBSD01} with File ID: VY022475.D met requirements for all samples except for Chloroethane[131%], failing high but no positive hit in associated samples therefore no corrective action taken. The RPD for {VY0530SBSD01} with File ID: VY022475.D met criteria except for Chloromethane[26%], Due to difference in results of BS- BSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements. |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: The Holding Times were not met for samples B27-SOIL-SAMPLE and B28-SOIL-SAMPLE, as samples activated out of hold. |  |  |  |

ADDITIONAL COMMENTS:

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date