| ORDER ID: Q2198 | MATRIX: Solid, Water |
| --- | --- |
| METHOD: 6010D,7470A,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution (OK-01-060325L) met criteria for all samples except for Calcium, Chromium, Copper, Manganese due to sample matrix interference. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (B-202-GW01MS) analysis met criteria for all elements except for Antimony and Manganese due to Chemical Interference during Digestion process. The Matrix Spike (OK-01-060325MS) analysis met criteria for all elements except for Antimony, Barium, Beryllium,Potassium, Sodium and Zinc due to Chemical Interference during Digestion Process and for Mercury due to sample matrix interference. The Matrix Spike Duplicate (B-202-GW01MSD) analysis met criteria for all elements except for Antimony and Manganese due to Chemical Interference during Digestion Process. The Matrix Spike Duplicate (OK-01-060325MSD) analysis met criteria for all elements except for Antimony, Barium, Beryllium, Chromium, Copper, Potassium, Silver and Sodium due to Chemical Interference during Digestion process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (OK-01-060325DUP) analysis met criteria for all samples except for Cadmium and Sodium due to sample matrix interference. The Duplicate (OK-01-060325MSD) analysis met criteria for all samples except for Potassium due to Chemical Interference during Digestion Process. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: The Post Digest Spike (Ok-01-060325A) analysis met criteria for all elements except for Barium, Chromium, Copper,  Potassium and Sodium  due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis , matrix has suppression effect during addition of spike.

The Post Digest Spike (B-202-GW01A) analysis met criteria for all elements except for Antimony and Manganese  due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis , matrix has suppression effect during addition of spike.

In analytical sequence LB136011, The % recovery was outside of acceptance limit for Iron, Potassium and Sodium of CCV08 but, no any sample associated under this CCV.

In analytical sequence LB136052, The % recovery was outside of acceptance limit for Arsenic, Iron, Potassium, Selenium, Sodium, Thallium, Zinc of CCV06 and for Arsenic and Selenium of CCV07 but, no any samples associated under this CCVs.

_______________________________________				__________________________

QA REVIEW								Date