| ORDER ID: Q2268 | MATRIX: Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | Compound #09 removed from 80 PPM.Compound #09 removed from 80 PPM. The Continuous Calibration File ID BF142748.D met the requirements except for 2,4-Dinitrophenol . But associated samples have not positive hit for this compound therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2268-04MS} with File ID: BF142735.D recoveries met the requirements for all compounds except for 2,3,4,6-Tetrachlorophenol[89%] and Hexachloroethane[211%] due to matrix interference. The MSD {Q2268-05MSD} with File ID: BF142736.D recoveries met the acceptable requirements except for 2,3,4,6-Tetrachlorophenol[85%] and Hexachloroethane[192%] due to matrix interference. The Blank Spike met requirements for all samples . The RPD for {Q2268-05MSD} with File ID: BF142736.D met criteria except for Naphthalene[42%] due to difference in MS and MSD concentrations. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for MW-6-20250605, MW-3-20250605. Due to high concentration of compounds, these samples required dilution. Therefore, samples were reanalyzed with dilution and reported. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW-2-20250605, MW-2-20250605-A, MW-6-20250605 and MW-3-20250605 were diluted due to high concentrations.

The SDG Q2268 is logged for SVOCMS Group1. Lab is not certified for 2,4-Dimethylphenol compound for 8270E method.

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date