| ORDER ID: Q2268 | MATRIX: Water |
| --- | --- |
| METHOD: 8260-Low |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID VN086967.D met the requirements except for Bromoform . But associated samples have not positive hit for these compounds therefore no corrective action was taken. The Continuous Calibration File ID VN087017.D met the requirements except for Carbon Disulfide . But associated samples have not positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2268-04MS} with File ID: VN086992.D recoveries met the requirements for all compounds except for Ethyl Benzene[240%], m/p-Xylenes[180%] due to matrix interference. The MSD {Q2268-05MSD} with File ID: VN086993.D recoveries met the acceptable requirements except for 1,2-Dibromoethane[124%], 1,2-Dichlorobenzene[120%], 1,3-Dichlorobenzene[123%], 1,4-Dichlorobenzene[122%], Benzene[240%], Chlorobenzene[121%], Chloroform[129%], Cyclohexane[123%], Ethyl Benzene[580%], Isopropylbenzene[176%], m/p-Xylenes[410%], Methyl tert-butyl Ether[138%], Methylcyclohexane[138%] and Toluene[126%] due to matrix interference. The Blank Spike for {VN0612WBS01} with File ID: VN086970.D met requirements for all samples except for 1,1,1-Trichloroethane[109%], 1,1,2-Trichloroethane[114%], 1,2-Dibromoethane[111%], 1,2-Dichlorobenzene[110%], 1,2-Dichloropropane[112%], 1,4-Dichlorobenzene[110%], Bromoform[115%], Dibromochloromethane[112%] and t-1,3-Dichloropropene[111%] but associated samples have positive hit for these compounds therefore no corrective action was taken. The Blank Spike for {VN0613WBS01} with File ID: VN087000.D met requirements for all samples except for 1,1-Dichloroethene[111%], Bromoform[114%], Carbon disulfide[117%] and o-Xylene[110%] but associated samples have positive hit for these compounds therefore no corrective action was taken.. The Blank Spike Duplicate met requirements for all samples . The RPD for {Q2268-05MSD} with File ID: VN086993.D met criteria except for 1,1,2-Trichlorotrifluoroethane[21%], Benzene[67%], Bromomethane[31%], Cyclohexane[40%], Dichlorodifluoromethane[21%], Ethyl Benzene[83%], Isopropylbenzene[41%], m/p-Xylenes[78%], Methylcyclohexane[49%] and Toluene[21%] due to difference in MS and MSD concentrations. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW-2-20250605, MW-2-20250605-A, MW-6-20250605 and MW-3-20250605 were diluted due to high concentrations.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date