| ORDER ID: Q2299 | MATRIX: Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for RW10A-MW01S-20250609 [Terphenyl-d14 - 136%], TT158I1-20250610 [Terphenyl-d14 - 142%], DUP01-20250610 [Terphenyl-d14 - 182%], RE131D2-20250610DL [Terphenyl-d14 - 151%], DUP02-20250610 [Terphenyl-d14 - 212%], DUP02-20250610DL [Nitrobenzene-d5 - 49% and Terphenyl-d14 - 288%],These failure surrogates are not associated with the client list, as per criteria affected surrogates were passing, therefore nocorrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2299-03MS} with File ID: BN037246.D recoveries met the requirements for all compounds except for 1,4-Dioxane[-50%],due to matrix interference. The MSD {Q2299-04MSD} with File ID: BN037247.D recoveries met the acceptable requirements except for 1,4-Dioxane[0%],due to matrix interference. The Blank Spike met requirements for all samples . The RPD for {Q2299-04MSD} with File ID: BN037247.D met criteria except for 1,4-Dioxane[200%],due to difference in results of MS and MSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The sample # TT191D1-20250609MS and TT191D1-20250609MSD is failing for 1,4-Dioxane and the original sample(TT191D1-20250609) is reported with M flag for this compounds.

Samples RE131D2-20250610, DUP02-20250610 and RE134D3-20250610 were diluted due to high concentrations.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date