| ORDER ID: Q2316 | MATRIX: Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike met requirements for all samples. The Blank Spike Duplicate met requirements for all samples. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for PB168476BS , Not associated, therefor no further action was taken. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).”

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date