| ORDER ID: Q2342 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS recoveries for {Q2347-01MS} with File ID: FE054490.D met requirements for all samples except for Aliphatic [n-Decane(C10)-142%], due to matrix interference. The MSD {Q2347-01MSD} with File ID: FE054491.D recoveries met requirements for all samples except for Aliphatic [n-Decane(C10)- 141%], due to matrix interference. The Blank Spike for met requirements The Blank Spike Duplicate met requirements The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date