| ORDER ID: Q2364 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (GAV1BMS) analysis met criteria for all parameters except for Antimony, Barium, Manganese, Sodium due to matrix interference. The Matrix Spike Duplicate (GAV1BMSD) analysis met criteria for all parameters except for Antimony, Barium, Sodium due to matrix interference. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (GAV1BDUP) analysis met criteria for all parameters except for Cadmium, Calcium, Cobalt, Copper, Lead, Manganese, Zinc due to matrix interference. The Duplicate (GAV1BMSD) analysis met criteria for all parameters except for Copper due to matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: The Post Digest Spike (GAV1BA) analysis met criteria for all parameters except for Manganese due to unknown chemical interferences of matrix with the addition of spike amount after digestion and before analysis , matrix has suppression effect during addition of spike.

In analytical sequence LB136273, the concentration was outside of acceptance limit for Chromium of CCB05 which is not associated to any sample of this project.

_______________________________________				__________________________

QA REVIEW								Date