| ORDER ID: Q2413 | MATRIX: Solid |
| --- | --- |
| METHOD: 8151A/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2405-01MS} with File ID: PS030842.D recoveries met the requirements for all compounds except for [Dinoseb(1)-8%], [Dinoseb(2)-2%] due to matrix interference . The MSD {Q2405-01MSD} with File ID: PS030843.D recoveries met the acceptable requirements except for [Dinoseb(1)-9%], [Dinoseb(2)-8%] due to matrix interference . The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date