| ORDER ID: Q2425 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution (COMP-3L) met criteria for all samples except for Aluminum, Iron and Manganese due to unknow sample matrix interference. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (COMP-3MS) analysis met criteria for all parameters except for Antimony, Boron, Molybdenum, Selenium, Silver and Zinc due to Chemical Interference during Digestion process. The Matrix Spike Duplicate (COMP-3MSD) analysis met criteria for all parameters except for Antimony, Molybdenum, Potassium, Silver, Vanadium and Zinc due to Chemical Interference during Digestion process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: This Data Package has been Revised due to Boron and Molybdenum parameters added in parameter list as per Client Request.

FAX and Hard copy Data Not Match Due to at time of FAX analysis Boron and Molybdenum parameter Recovery out side limit  for ICV in sequence so its Re-analyse in another sequence , so in Hard copy correct data reported.

_______________________________________				__________________________

QA REVIEW								Date