| ORDER ID: Q2458 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | Compound #09 removed from 80 PPM.Compound #09 removed from 80 PPM.Compound #09 removed from 80 PPM.Compound #09 removed from 80 PPM. The Continuous Calibration File ID BF142941.D met the requirements except for 4-Nitrophenol ,is marginally biased low therefore no Corrective action was taken. The Continuous Calibration File ID BF142964.D met the requirements except for Hexachlorocyclopentadiene is marginally biased low therefore no Corrective action was taken. The Continuous Calibration File ID BF143025.D met the requirements except for 2,4-Dinitrophenol,4-Nitrophenol and Pentachlorophenol,are biased high therefore no Corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB168674BS} with File ID: BF142991.D met requirements for all samples except for Butylbenzylphthalate[106%],is failing high but no positive hit in associate samples therefore no corrective action taken. The Blank Spike for {PB168716BS} with File ID: BF143027.D met requirements for all samples except for Butylbenzylphthalate[106%],is failing high but no positive hit in associate samples therefore no corrective action taken. The Blank Spike Duplicate met requirements for all samples .The RPD for {PB168716BSD} with File ID: BF143028.D met criteria except for 3,3-Dichlorobenzidine[35%], 4-Chloroaniline[22%],due to difference in results of BS and BSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples TP-63 was diluted dueto dirty, concentrated and viscous matrix.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date