| ORDER ID: Q2480 | MATRIX: Solid |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82W063025S.M) for Methylene Chloride passing on Quadratic Regression. The Continuous Calibration File ID VX046933.D met the requirements except for 2-Hexanone and Methyl Acetate are failing high but no positive hit in associate sample therefore no corrective action taken. The Continuous Calibration File ID VX046961.D met the requirements except for 1,2-Dibromo-3-Chloropropane,2-Butanone,2-Hexanone,4-Methyl-2-Pentanone,Methyl Acetate and Methyl tert-butyl Ether are failing high but no positive hit in associate sample therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for GPX2 [4-Bromofluorobenzene - 226%, Toluene-d8 - 163%], GPX6 [4-Bromofluorobenzene - 286%, Toluene-d8 - 219%], GPX7 [4-Bromofluorobenzene - 201% and Toluene-d8 - 209%] Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported, while for GPX3RE [Toluene-d8 - 72%] sample was reanalyzed to confirm the failure and reported. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {VW0709SBS01} with File ID: VW031769.D met requirements for all samples except for cis-1,2-Dichloroethene[124%], Methylene Chloride[176%] are failing high but no positive hit in associate sample therefore no corrective action taken. The Blank Spike for {VX0710MBS01} with File ID: VX046953.D met requirements for all samples except for Bromomethane[55%] failing biased low but associate CCAL passing for failing analyte and also sample having no positive hit for Bromomethane therefore no corrective action taken. The Blank Spike for {VX0714MBS01} with File ID: VX046964.D met requirements for all samples except for 2-Butanone[135%], 2-Hexanone[141%] and 4-Methyl-2-Pentanone[136%] are failing high but no positive hit in associate sample therefore no corrective action taken. The Blank Spike for {VY0707SBS01} with File ID: VY022948.D met requirements for all samples except for Methylene Chloride[142%] is failing high and associate sample having hit of Methylene Chloride but below CRQL therefore no corrective action taken. The Blank Spike for {VY0708SBS01} with File ID: VY022970.D met requirements for all samples except for Methylene Chloride[149%] is failing high and associate sample having positive hit of Methylene Chloride therefore as corrective action lab reanalyzed and reported both run. The Blank Spike Duplicate met requirements for all samples . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for GPX3RE, sample was reanalyzed to confirm the failure and reported. |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples GPX5, GPX8 were directly analyzed in methanol as per past history of the samples.

Samples GPX2, GPX4, GPX6 and GPX7 were diluted due to high concentrations.

The Sample #GPX1, GPX3, GPX5, GPX6ME , GPX7ME, GPX8 have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date