| ORDER ID: Q2481 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8270E/3510/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for, CC0627-CLOXPL [2,4,6-Tribromophenol - 2058%, 2-Fluorobiphenyl - 2551%, 2-Fluorophenol - 1%, Nitrobenzene-d5 - 15%, Phenol-d6 - 0%, Terphenyl-d14 - 15%], CC0625-OXBL [2-Fluorophenol - 7%, Nitrobenzene-d5 - 8%, Phenol-d6 - 4%, Terphenyl-d14 - 11%], CC0627-AOXL [2,4,6-Tribromophenol - 1262%, 2-Fluorobiphenyl - 1236%, Phenol-d6 - 2%, Terphenyl-d14 - 3535%], CC0625-NL [2,4,6-Tribromophenol - 7%, 2-Fluorobiphenyl - 10%, 2-Fluorophenol - 8%, Nitrobenzene-d5 - 9%, Phenol-d6 - 8%, Terphenyl-d14 - 12%], CC0267-OXPL [2,4,6-Tribromophenol - 3918%, 2-Fluorobiphenyl - 4832%, 2-Fluorophenol - 2%, Nitrobenzene-d5 - 8%, Phenol-d6 - 0%, Terphenyl-d14 - 11%], CC0627-OXL [2,4,6-Tribromophenol - 8108%, 2-Fluorobiphenyl - 10034%, 2-Fluorophenol - 0%, Nitrobenzene-d5 - 48%, Phenol-d6 - 0%], CC0627-CLOXAL [2,4,6-Tribromophenol - 0%, 2-Fluorophenol - 0%, Nitrobenzene-d5 - 11%, Phenol-d6 - 0%, Terphenyl-d14 - 0%], CC0627-BL [2,4,6-Tribromophenol - 2%, 2-Fluorobiphenyl - 10%, 2-Fluorophenol - 1%, Nitrobenzene-d5 - 8%, Phenol-d6 - 2%, Terphenyl-d14 - 10%], CC0627-SFBL [2,4,6-Tribromophenol - 25%, 2-Fluorobiphenyl - 37%, Nitrobenzene-d5 - 34% and Terphenyl-d14 - 27%]. As evidenced by the chromatograms, recoveries of the surrogates were affected by the samples matrix. Therefor no further corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike met requirements for all compounds. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas were met for all analysis except for CC0627-CLOXPL, CC0625-OXBL, CC0627-AOXL, CC0625-NL, CC0267-OXPL, CC0627-OXL, CC0627-CLOXAL and CC0627-BL. As evidenced by the chromatograms internal standard were affected by the samples matrix. Therefor no further corrective action was taken. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: The Holding Times were met for all analysis. |  |  |  |

ADDITIONAL COMMENTS:

Samples CC0627-CLOXPL, CC0625-OXBL, CC0625-NL, CC0267-OXPL, CC0627-OXL, CC0627-CLOXAL and CC0627-BL all samples were Chemical treated, Therefor samples were analyzed with direct 10X dilution factor.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

All samples used limited volume as samples are not regular environmental samples it is chemical treated samples.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date