| ORDER ID: Q2481 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8151A/3510/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for CC0627-CLOXPL [2,4-DCAA(1)49%], CC0625-OXBL [2,4-DCAA(1)252%, 2,4-DCAA(2)172%], CC0625-NL [2,4-DCAA(1)37%, 2,4-DCAA(2)453%], CC0267-OXPL [2,4-DCAA(1)7%, 2,4-DCAA(2)254%], CC0627-OXL [2,4-DCAA(1)2153%] and CC0627-BL [2,4-DCAA(2)626%]. As evidenced by the chromatograms, recoveries of the surrogates were affected by the samples matrix. Therefor no further corrective action was taken. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2641-02MS} with File ID: PS031234.D recoveries met the requirements for all compounds except for [2,4,5-TP(Silvex)(1)155% - 2,4,5-TP(Silvex)(2)153%] and [2,4-D(1)161% - 2,4-D(2)172%] due to matrix interference.. The MSD {Q2641-02MSD} with File ID: PS031235.D recoveries met the requirements for all compounds except for [2,4,5-TP(Silvex)(1)148% - 2,4,5-TP(Silvex)(2)145%] and [2,4-D(1)152% - 2,4-D(2)162%] due to matrix interference. The Blank Spike met requirements for all compounds. The RPD were met for all analysis. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. The Holding Times were met for all analysis. |  |  |  |

ADDITIONAL COMMENTS:

Samples CC0627-CLOXPL, CC0625-OXBL, CC0625-NL, CC0267-OXPL, CC0627-OXL and CC0627-BL were diluted due to non environmental chemical treated samples are received and having very bad matrix.

The fax and hardcopy is not matching for CC0627-AL and CC0627-OXL, as fax samples 	 analyzed in seq PS072825 where Ending I.BLK and CCAL missing,  as corrective action lab reanalyzed these samples in seq PS073025 and hard copy reported from second analysis. The above sample original run is reported as screening data in miscellaneous data.

The temperature of the samples at the time of receipt was 21.0°C.

Less volume was taken at the time of extraction due to these samples were not regularenvironmental samples, its chemical treated sample

_______________________________________				__________________________

QA REVIEW								Date