| ORDER ID: Q2481 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8081B/3510/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for CC0627-CLOXPL [Tetrachloro-m-xylene(1)655%, Tetrachloro-m-xylene(2)386%], CC0625-OXBL [Decachlorobiphenyl(1)998%, Decachlorobiphenyl(2)794%, Tetrachloro-m-xylene(1)432%, Tetrachloro-m-xylene(2)44%], CC0625-NL [Decachlorobiphenyl(1)234%], CC0267-OXPL [Tetrachloro-m-xylene(2)47%], CC0627-OXL [Tetrachloro-m-xylene(1)1168%], CC0627-CLOXAL [Decachlorobiphenyl(2)49%, Tetrachloro-m-xylene(1)82241%, Tetrachloro-m-xylene(2)1375%] and CC0627-BL [Tetrachloro-m-xylene(1)573%]. As evidenced by the chromatograms, recoveries of the surrogates were affected by the samples matrix. Therefor no further corrective action was taken. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS recoveries met the requirements for all compounds. The MSD recoveries met the requirements for all compounds. The Blank Spike met requirements for all compounds. The RPD were met for all analysis. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. The Holding Times were met for all samples except for CC0627-AL, CC0627-CLOXPL, CC0625-OXBL, CC0627-AOXL, CC0625-NL, CC0267-OXPL, CC0627-OXL, CC0627-CLOXAL, CC0627-BL and CC0627-SFBL. |  |  |  |

ADDITIONAL COMMENTS:

Samples CC0627-CLOXPL, CC0625-OXBL, CC0625-NL, CC0267-OXPL, CC0627-OXL, CC0627-CLOXAL and CC0627-BL were diluted due to non environmental chemical treated samples are received and having very bad matrix.

The temperature of the samples at the time of receipt was 21.0°C.

Less volume was taken at the time of extraction due to these samples were not regular environmental samples, its chemical treated sample

_______________________________________				__________________________

QA REVIEW								Date