| ORDER ID: Q2481 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8260D/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for CC0625-OXBL [Dibromofluoromethane - 59%], CC0627-AOXL [Toluene-d8 - 116%], CC0627-BL [Dibromofluoromethane - 55%] and CC0627-SFBL [Dibromofluoromethane - 37%] due to bad matrix and also samples have limited volume therefore no corrective action taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas were met for all analysis except for CC0627-BL, CC0627-SFBL due to bad matrix and also samples have limited volume therefore no corrective action taken. |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: The Holding Times were met for all analysis. |  |  |  |

ADDITIONAL COMMENTS:

Samples CC0627-AL, CC0627-CLOXPL, CC0625-OXBL, CC0627-AOXL, CC0625-NL, CC0267-OXPL, CC0627-OXL, CC0627-CLOXAL, CC0627-BL and CC0627-SFBL were diluted due to bad matrix and also samples have limited volume not allowing any sample to be run straight.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date