| ORDER ID: Q2484 | MATRIX: Solid |
| --- | --- |
| METHOD: 8081B/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for TP-58 [Decachlorobiphenyl(1)193% ]. As per method one surrogate allowed to fail to meet the criteria per column. No further corrective action was taken. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS recoveries met the requirements for all compounds due to matrix interference. The MSD recoveries met the acceptable requirements due to matrix interference. The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date