| ORDER ID: Q2487 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution met criteria for all compounds. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (WC-11MS) analysis met criteria for all compounds except for Antimony, Selenium, Silver, Thallium, Vanadium due to chemical interference during digestion Process. The Matrix Spike Duplicate (WC-11MSD) analysis met criteria for all compounds except for Antimony, Potassium, Selenium, Silver, Thallium, Vanadium due to chemical interference during digestion process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (WC-11DUP) analysis met criteria for all compounds except for Cadmium due to sample matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: The Post Digest Spike (WC-11A) analysis met criteria for all compounds except for Silver due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

In analytical sequence LB136390, The % recovery was outside of acceptance limit for Thallium of CCV05, but no any sample associated under this CCV.

In analytical sequence LB136390, The % recovery was outside of acceptance limit for Potassium of CCV06, but no any sample associated under this CCV.

In analytical sequence LB136390, The % recovery was outside of acceptance limit for Thallium of CCV06., but no any sample associated under this CCV.

In analytical sequence LB136434, The % recovery was outside of acceptance limit for Potassium of ICV01, but no any sample associated under this ICV.

In analytical sequence LB136434, The % recovery was outside of acceptance limit for Silver of ICV01, but no any sample associated under this ICV.

_______________________________________				__________________________

QA REVIEW								Date