| ORDER ID: Q2487 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BP070325.M) for 2-Nitrophenol,2-Nitroaniline,2,4-Dinitrophenol,4,6-Dinitro-2-methylphenol,Butylbenzylphthalate,Di-n-octyl phthalate these compounds are passing on LinearRegression. The Continuous Calibration File ID BF143025.D met the requirements except for 2,4-Dinitrophenol,4-Nitrophenol and Pentachlorophenol,are biased failing high but no positive hit in associate samples therefore no corrective action taken. The Continuous Calibration File ID BF143048.D met the requirements except for 2,4-Dinitrophenol and Pentachlorophenol,are biased failing high but no positive hit in associate samples therefore no corrective action taken. The Continuous Calibration File ID BP025082.D met the requirements except for Hexachlorocyclopentadiene,2,4-Dinitrophenol,4,6-Dinitro-2-methylphenol and 3,3-Dichlorobenzidine,are failing high but no positive hit in associate samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike met requirements for all samples . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for G3(0-6), G1(0-6) and G1(6-12),due to the presence of non-targeted hydrocarbons, which can be observed from the abnormal chromatogram therefore no corerective action was taken. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples G4(0-6), G3(0-6), G2(0-6), G1(0-6) and G1(6-12) were diluted due to viscous, dirty and concentrated matrix.

Samples G2(0-6), G2(0-6)DL were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date