| ORDER ID: Q2515 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution (WC-1L) met criteria for all compounds except for Iron, Magnesium and Manganese due to sample matrix interference. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (WC-1MS) analysis met criteria for all compounds except for Antimony, Beryllium, Cobalt, Copper, Selenium, Silver, Sodium and Vanadium due to Chemical Interference during Digestion Process. The Matrix Spike Duplicate (WC-1MSD) analysis met criteria for all compounds except for Antimony, Cobalt, Copper, Selenium, Silver, Sodium and Vanadium due to Chemical Interference during Digestion Process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (WC-1DUP) analysis met criteria for all compounds except for Lead due to sample matrix interference. The Duplicate (WC-1MSD) analysis met criteria for all compounds except for Manganese and Vanadium due to Chemical Interference during Digestion Process. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: The Post Digest Spike (WC-1A) analysis met criteria for all compounds except for Antimony, Copper, Selenium, Silver, Sodium and Vanadium  due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

In analytical sequence LB136434, The Results was outside of acceptance limit for Silver of CCB08 but no any sample associated under this CCB.

_______________________________________				__________________________

QA REVIEW								Date