[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

CDM Smith
Project Name: South River WM Replacement
Project # N/A
Order ID # Q2529
Test Name: SVOC-TCL BNA -20
A. Number of Samples and Date of Receipt:

10 Solid samples were received on 07/08/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-TCL BNA -20. This data package contains results for SVOC-TCL BNA -20.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The samples were analyzed on instrument BNA_M using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The samples were analyzed on instrument BNA_P using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-TCL BNA -20 was based on method 8270E and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the requirements for all compounds.
The RPD for {Q2517-01MSD} with File ID: BF143061.D met criteria except for 3-Nitroaniline[21%], 4-Chloroaniline[22%],due to difference in results of MS and MSD.
The Blank Spike for {PB168767BS} with File ID: BP025084.D met requirements for all compounds except for 2,4,5-Trichlorophenol[106%], 2,4,6-Trichlorophenol[106%], 2,6-Dinitrotoluene[106%], 4-Nitroaniline[106%], Acenaphthylene[106%], Anthracene[106%], Benzo(a)pyrene[112%], Carbazole[100%], Di-n-butylphthalate[106%], Hexachlorobenzene[100%], Hexachlorocyclopentadiene[209%], Indeno(1,2,3-cd)pyrene[106%], Isophorone[100%], N-Nitrosodiphenylamine[100%] and Pentachlorophenol[106%],are marginally biased high, however they are passing in the Continious calibration,therefore no corrective action was taken.
The Blank analysis did not indicate the presence of lab contamination.
The % RSD is greater than 20% in the Initial Calibration (8270-BM070825.M) for 2,4-Dinitrophenol, 2,4-Dinitrotoluene,4-Nitroaniline,4,6-Dinitro-2-methylphenol,these compounds are passing on Linear Regression.

The % RSD is greater than 20% in the Initial Calibration (8270-BP070325.M) for 2-Nitrophenol,2-Nitroaniline,2,4-Dinitrophenol,2,4-Dinitrotoluene,4,6-Dinitro-2-methylphenol,Butylbenzylphthalate,Di-n-octyl phthalate these compounds are passing on Linear Regression.

The Continuous Calibration File ID BF143048.D met the requirements except for 2,4-Dinitrophenol and Pentachlorophenol,are biased failing high  but no positive hit in associate samples therefore no corrective action taken.

The Continuous Calibration File ID BP025155.D met the requirements except for 2,4-Dinitrophenol,2-Nitrophenol,3,3-Dichlorobenzidine,4,6-Dinitro-2-methylphenol,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol ,are biased failing high and ,Bis(2-ethylhexyl)phthalate is marginally biased high but no positive hit in associated samples therefore no corrective action taken.
The Tuning criteria met requirements.  
Samples TP-102 was diluted due to to viscous, dirty and concentrated matrix.
E. Additional Comments:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


