284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922

CASE NARRATIVE

### Alliance Technical Group, LLC - Newark

### Project Name: NJ Drinking Water PT

### Project # N/A

### Order ID # Q2552

Test Name: Turbidity

A. Number of Samples and Date of Receipt:

2 Water samples were received on 07/09/2025.

B. Parameters:

According to the Chain of Custody document, the following analyses were requested:

Turbidity. This data package contains results for Turbidity.

C. Analytical Techniques:

The analysis of Turbidity was based on method 180.1 and The analysis of Turbidity was based on method SM2130 B.

D. QA/ QC Samples:

The Holding Times were met for all samples except for WS0725-PT-TURB-WS of Turbidity and for WS0725-PT-TURB-WS of Turbidity as samples were receive out of holding time.

The Blank Spike met requirements for all compounds.

The Duplicate analysis met criteria for all compounds.

The Blank analysis did not indicate the presence of lab contamination.

The Calibration met the requirements.

E. Additional Comments:

The time of sampling was not listed in the COC.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________