| ORDER ID: Q2580 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270-Modified/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for WC-URBAN-FILL-B2DL [2,4,6-Tribromophenol - 0%, 2-Fluorobiphenyl - 0%, 2-Fluorophenol - 0%, 2-Methylnaphthalene-d10 - 0%, Fluoranthene-d10 - 0%, Nitrobenzene-d5 - 0%, Phenol-d6 - 0% and Terphenyl-d14 - 0%],DMC recovery do not apply for samples that have been diluted therefore no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2580-02MS} with File ID: BN037491.D recoveries met the requirements for all compounds except for 2-Methylnaphthalene[-71%], Acenaphthene[-355%], Acenaphthylene[-71%], Anthracene[-709%], Benzo(a)anthracene[-709%], Benzo(a)pyrene[-709%], Benzo(b)fluoranthene[-709%], Benzo(g,h,i)perylene[-709%], Benzo(k)fluoranthene[-213%], Chrysene[-709%], Dibenz(a,h)anthracene[-142%], Fluoranthene[-1418%], Fluorene[-213%], Indeno(1,2,3-cd)pyrene[-709%], Naphthalene[-496%], Phenanthrene[-2837%] and Pyrene[-2128%],These compounds did not meet the NJDKQP criteria and in-house criteria ,due to matrix interference. The MSD {Q2580-02MSD} with File ID: BN037492.D recoveries met the requirements for all compounds except for 2-Methylnaphthalene[-71%], Acenaphthene[-284%], Anthracene[-709%], Benzo(a)anthracene[-1418%], Benzo(a)pyrene[0%], Benzo(b)fluoranthene[-709%], Benzo(g,h,i)perylene[-709%], Benzo(k)fluoranthene[142%], Chrysene[0%], Dibenz(a,h)anthracene[-71%], Fluoranthene[709%], Fluorene[-142%], Indeno(1,2,3-cd)pyrene[0%], Naphthalene[-496%], Phenanthrene[-2128%] and Pyrene[-2837%],these compounds did not meet the NJDKQP criteria and in-house criteria ,due to matrix interference. The Blank Spike met requirements for all compounds. The RPD for {Q2580-02MSD} with File ID: BN037492.D met criteria except for Benzo(a)anthracene[67%], Benzo(a)pyrene[200%], Benzo(k)fluoranthene[1000%], Chrysene[200%], Dibenz(a,h)anthracene[67%], Fluoranthene[600%], Fluorene[40%], Indeno(1,2,3-cd)pyrene[200%],these compounds did not meet the NJDKQP criteria and in-house criteria While Acenaphthene[22%],Phenanthrene[29%] and Pyrene[29%].these compounds met the NJDKQP criteria but did not meet the in-house criteria,due to difference in results of MS and MSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas were met for all analysis except for WC-URBAN-FILL-B2MS and WC-URBAN-FILL-B2MSD, due to concentrated matrix therefore no corrective action was taken. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples WC-URBAN-FILL-B2, WC-URBAN-FILL-B2MS and WC-URBAN-FILL-B2MSD were diluted due to bad matrix.

Sample WC-URBAN-FILL-B2 was diluted due to high concentration.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date