| ORDER ID: Q2600 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2600-05MS} with File ID: FC069465.D recoveries met the requirements for all compounds except for Aliphatic [ n-Docosane (C22)- 155%], [n-Hexacosane (C26) - 328%], [n-Octacosane (C28) - 222%], [n-Dotriacontane (C32) -165%] due to matrix interfernce. & for [Naphthalene (C11.7)- 16%, 2-methylnaphthalene (C12.89)- 5%], these analytes are only being monitoring in aliphatic fraction.The MS {Q2600-05MS} with File ID: FD049555.D recoveries met the requirements for all compounds except for Aromatic [Benzo[a]anthracene (C26.37)- 148%], [Chrysene (C27.41) - 291%], [Bnezo[k]fluoranthene (C30.14) - 264%], [Dibenz[a,h]anthracene (C30.36) -222%], [Benzo[a]pyrene (C31.34) - 142%]due to matrix interference. The MS {Q2600-05MSD} with File ID: FC069466.D recoveries met the requirements for all compounds except for Aliphatic [n-Nonane (C9) -39%], [ n-Docosane (C22)- 143%], [n-Hexacosane (C26) - 214%], [n-Octacosane (C28) - 276%], [n-Dotriacontane (C32) -146%] due to matrix interfernce. & for [Naphthalene (C11.7)- 16%, 2-methylnaphthalene (C12.89)- 5%], these analytes are only being monitoring in aliphatic fraction. The MS {Q2600-05MS} with File ID: FD049555.D recoveries met the requirements for all compounds except for Aromatic [Benzo[a]anthracene (C26.37)- 143%], [Chrysene (C27.41) - 265%], [Bnezo[k]fluoranthene (C30.14) - 262%], [Dibenz[a,h]anthracene (C30.36) -226%], [Benzo[a]pyrene (C31.34) - 156%]due to matrix interference. The Blank Spike met requirements for all compounds. The Blank Spike Duplicate met requirements for all compounds. The RPD were met for all analysis. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples TRENCH, STOCK-PILE and END-OF-TRENCH were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date