| ORDER ID: Q2600 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF071525.M) for 2-Nitrophenol, 2-Nitroaniline, 2,6-Dinitrotoluene, 2,4-Dinitrotoluene,Butylbenzylphthalate, Bis(2-ethylhexyl)phthalate, and Di-n-octyl phthalate thesecompounds are passing on Linear Regression and 2,4-Dinitrophenol,4,6-Dinitro-2-methylphenol, these Compounds are passing on Qudratic regression. The % RSD is greater than 20% in the Initial Calibration (8270-BP070325.M) for 2-Nitrophenol,2-Nitroaniline,2,4-Dinitrophenol,2,4-Dinitrotoluene,4,6-Dinitro-2-methylphenol,Butylbenzylphthalate,Di-n-octyl phthalate these compounds are passing onLinear Regression. The Continuous Calibration File ID BP025155.D met the requirements except for 2,4-Dinitrophenol,2-Nitrophenol,3,3-Dichlorobenzidine,4,6-Dinitro-2- methylphenol,Bis(2-ethylhexyl)phthalate,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol are biased high therefore no corrective action was taken. The Continuous Calibration File ID BP025187.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol, are biased high,and no positive hit inassociated sampls, therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB168854BS} with File ID: BF143119.D met requirements for all compounds except for 4-Nitroaniline[106%],is failing high but no positive hit in associate samples therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date