| ORDER ID: Q2645 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% for certain compounds in the Initial Calibration (Method 8270-BF071725.M) for 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol these Compound is passing on Linear Regression. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for RW5B-CARBON-20250716 [2,4,6-Tribromophenol - 3%, 2-Fluorobiphenyl - 7%, 2-Fluorophenol - 2%, Nitrobenzene-d5 - 9%, Phenol-d6 - 3% and Terphenyl-d14 - 4%]. The similar failure in the Pest / PCB sample was also observed. The surrogates in the re-extracted Pest/PCB sample failed again, confirming the matrix interference. Therefore, the sample RW5B-CARBON-20250716 was not re-extracted and also not re-analyzed, as the reanalysis would result in the same recoveries of the surrogates. Therefor no further corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2649-09MS} with File ID: BF143219.D recoveries met the requirements for all compounds except for 4-Chloroaniline[13%]. Recovery failed due to matrix interference. Therefor no further corrective action was taken. The MSD {Q2649-09MSD} with File ID: BF143220.D recoveries met the requirements for all compounds except for 4-Chloroaniline[16%]. Recovery failed due to matrix interference. Therefor no further corrective action was taken. The Blank Spike met requirements for all compounds. The RPD for {Q2649-09MSD} with File ID: BF143220.D met criteria except for 4-Chloroaniline[21%]. RPD failed due to result difference between MS and MSD, therefor no further corrective action was taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The laboratory certifies that the all-electronic diskette deliverable exactly match the data

summary forms (i.e. Form Is).”

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

The soil samples results are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date