| ORDER ID: Q2651 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (MH 2-1MS) analysis met criteria for all compounds except for Antimony and Zinc due to Chemical Interference during Digestion Process. The Matrix Spike Duplicate (MH 2-1MSD) analysis met criteria for all compounds except for Antimony and Zinc due to Chemical Interference during Digestion Process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (MH 2-1DUP) analysis met criteria for all compounds except for Calcium and Magnesium due to sample matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

_______________________________________				__________________________

QA REVIEW								Date