| ORDER ID: Q2669 | MATRIX: Water |
| --- | --- |
| METHOD: 624.1 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for 001 WILLETS PT BLVD (JUNE) [1,2-Dichloroethane-d4 - 111%], 001 WILLETS PT BLVD (JUNE)DL [1,2-Dichloroethane-d4 - 119%], 002 35th Ave (JUNE) [1,2-Dichloroethane-d4 - 114%], 002 35th Ave (JUNE)DL [1,2-Dichloroethane-d4 - 115%] Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported while for VN0722WBS01 [1,2-Dichloroethane-d4 - 111%] failing marginally high therefore no corrective action taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike Duplicate for {VN0722WBSD01} with File ID: VN087391.D met requirements for all compounds except for m/p-Xylenes[80%], o-Xylene[77%] failing biased low , but passing in Blank spike therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples 001 WILLETS PT BLVD (JUNE), 002 35th Ave (JUNE) were diluted due to high concentrations.

The temperature of the samples at the time of receipt was 17.5°C, Lab notified this issue

to the client. See the communication in shipping Document section.

“As per method, MS/MSD is required to be performed with the sample analysis. However, Lab did not receive sufficient volume to perform the MS/MSD therefore MS/MSD were not performed for this project. However, Lab has performed LCS/LCSD instead.”

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date