| ORDER ID: Q2743 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2743-01MS} with File ID: FC069640.D recoveries met the requirements for all compounds except for Aliphatic [ n-Nonane (C9)- 38%], [ n-Tetracosane (C24)- 29%], due to matrix interference. & [Naphthalene (C11.7)- 1%, 2-methylnaphthalene (C12.89)- 0%], these analytes are only being monitoring in aliphatic fraction.The MS {Q2743-01MS} with File ID: FD049621.D recoveries met the requirements for all compounds except for Aromatic [Chrysene (C27.41)- 187%], [Bnezo[k]fluoranthene (C30.14)- 181%], [Dibenz[a,h]anthracene (C30.36)- 166%], [benzo[b]fluoranthene (C30.41)- 170%], [Indeno[1,2,3-cd]pyrene (C35.01- 175%]due to matrix interference. The MS {Q2743-01MSD} with File ID: FC069641.D recoveries met the requirements for all compounds except for Aliphatic [ n-Nonane (C9)- 37%], [ n-Tetracosane (C24)- 29%], due to matrix interference. & [Naphthalene (C11.7)- 1%, 2-methylnaphthalene (C12.89)- 0%], these analytes are only being monitoring in aliphatic fraction. The MS {Q2743-01MSD} with File ID: FD049622.D recoveries met the requirements for all compounds except for Aromatic [Chrysene (C27.41)- 186%], [Bnezo[k]fluoranthene (C30.14)- 180%], [Dibenz[a,h]anthracene (C30.36)- 165%], [benzo[b]fluoranthene (C30.41)- 170%], [Indeno[1,2,3-cd]pyrene (C35.01- 174%]due to matrix interference. The Blank Spike for {PB169209BS} with File ID: FC069636.D met requirements for all samples except for aliphatic [Naphthalene (C11.7)- 0%, 2-methylnaphthalene (C12.89)- 0%], these analytes compounds are only being monitoring in aliphatic friction. The Blank Spike Duplicate for {PB169209BSD} with File ID: FC069637.D met requirements for all samples except for aliphatic [Naphthalene (C11.7)- 0%, 2- methylnaphthalene (C12.89)- 0%], these analytes compounds are only being monitoring in aliphatic friction. The RPD were met for all analysis. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date