| ORDER ID: Q2747 | MATRIX: Solid |
| --- | --- |
| METHOD: 8151A/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. The Initial Calibration met the requirements. The Continuous Calibration met the requirements. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for OU4-TS-GRILLO-OG2-073125 [2,4-DCAA(2)24%], OU4-TS-GRILLO-OG2-073125RE [2 and4-DCAA(2)20%]the failure sample in surrogate was reanalyzed to confirm the failure as per method and reported in the data. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2753-01MS} with File ID: PS031365.D recoveries met the requirements for all compounds except for [Dalapon(1)35% - Dalapon(2)66%], [Dinoseb(1)7% - Dinoseb(2)7%]due to matrix interference. The MSD {Q2753-01MSD} with File ID: PS031366.D recoveries met the requirements for all compounds except for [Dalapon(1)25% - Dalapon(2)69%], [Dinoseb(1)7% - Dinoseb(2)7%]due to matrix interference. The Blank Spike met requirements for all compounds. The RPD for {Q2753-01MSD} with File ID: PS031366.D met criteria except for Dalapon[33%]due to difference in results of MS-MSD. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

As per special requirement for this project form-1 are reported in mg/kg.

The not QT review data is reported in the Miscellaneous.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date